# Helen Lake Tower Check List

## Name 	Dan			Date   	4-25-02			Time (UTC)   16:46

File transfer

#### Transferred files from C:\data to Zip 250 (E:\): 	 100-103, 105-113, 115a, 115l

## Deleted files in C:\data\backup (also empty Recycle Bin):      90-99

Moved transferred files to C:\data\backup	  x

###### Shutdown computer and restart		x

Transfer data from Micromet datalogger	  x

| ## Cylinders | ### Tank Pressure (psi) | ### Line Pressure | ### S/N (if tank change) |
| --- | --- | --- | --- |
| ## Nafion N2 | 1700 | 3 |  |
| Flux N2 | 1300 | 3 |  |
| CO2 Low | 1900 | 4 |  |
| CO2 High | 1300 | 4 |  |
| CO2 Ref. | 1900 | 3 |  |

| ## Flow | ### Value |
| --- | --- |
| ## Ref CO2 (45) | 42 |
| Flux N2 (45) | 45 |
| Nafion N2 (51) | 55/51 |
| Sample (51) | 51 |

| #### Miscellaneous |  |
| --- | --- |
| #### Pressure Gauge | #### 3 |
| #### Voltage (from power supply under bench) | #### 15 |
| Current (Amps) | 4 |
| Utility meter (Kilowatts) | 03161 |
| Temperature (shown on heater) | 52 |
| Weather | Overcast, morning snow |

Comments:     Replaced pump.  All of the snow was melted.  There was a dusting from

a morning snow shower.  There was no data for day 104 or day 114.  Upon further

inspection of the replaced pump, it was discovered that one of the diaphragms was

almost completely disintegrated.  Would you prefer I changed the diaphragms every visit?  I imagine little pieces of rubber go through the system when they break like this.

Check: pumps running (inside/outside), Licor on and running , UPS operative (under bench, light bulb icon in system tray), PC Anywhere running (small icon of computer in system tray), phone operative (dialtone), breakers not tripped (under bench and at utility meter), shed is dry, tarp covering workbench, heater set to 56, trash emptied, cylinders secured, cabinets and shed locked, rain gauges are clear, tower and guylines up and clear of tree branches, snow cleared at tower base