# Helen Lake Tower Check List

## Name 	Ron				Date   Feb 28,02				Time (UTC)17:00

File transfer

#### Transferred files from C:\data to Zip 250 (E:\): 	  0204922a-l to 0205023a-l

## Deleted files in C:\data\backup (also empty Recycle Bin):    Rename Backup2

Moved transferred files to C:\data\backup1	  x

###### Shutdown computer and restart		x

Transfer data from Micromet datalogger	  x

| ## Cylinders | ### Tank Pressure (psi) | ### Line Pressure | ### S/N (if tank change) |
| --- | --- | --- | --- |
| ## Nafion N2 | 1860 | 3 |  |
| Flux N2 | 1440 | 3 |  |
| CO2 Low | 1900 | 4 |  |
| CO2 High | 1350 | 3.5 |  |
| CO2 Ref. | 250 | 4 |  |

| ## Flow | ### Value |
| --- | --- |
| ## Ref CO2 (45) | 29/45 |
| Flux N2 (45) | 26/45 |
| Nafion N2 (51) | 51 |
| Sample (51) | 55 |

| #### Miscellaneous |  |
| --- | --- |
| #### Pressure Gauge | #### 4.4 |
| #### Voltage (from power supply under bench) | #### ~15 |
| Current (Amps) | ~5 |
| Utility meter (Kilowatts) | 0 2.5 4 7 6 |
| Temperature (shown on heater) | High teens |
| Weather | Overcast light snow |

Comments:     Everything was off=dead. Found GFI tripped. Reset GFI and Licor and pump started.  Restarted computer with SDM#16.  Computer collected data from loggers.  All looks good now.

Check: pumps running (inside/outside), Licor on and running , UPS operative (under bench, light bulb icon in system tray), PC Anywhere running (small icon of computer in system tray), phone operative (dialtone), breakers not tripped (under bench and at utility meter), shed is dry, tarp covering workbench, heater set to 56, trash emptied, cylinders secured, cabinets and shed locked, rain gauges are clear, tower and guylines up and clear of tree branches, snow cleared at tower base