# Helen Lake Tower Check List

## Name  Dan				Date   July 18 2005			Time (UTC) 19:47

File transfer

#### Transferred files from C:\data to Zip 250 (E:\): 	171-198

## Deleted files in C:\data\backup (also empty Recycle Bin):      160-170

Moved transferred files to C:\data\backup	  x

###### Shutdown computer and restart		x

Transfer data from Micromet datalogger	  x

| ## Cylinders | ### Tank Pressure (psi) | ### Line Pressure | ### S/N (if tank change) |
| --- | --- | --- | --- |
| ## Nafion N2 | 1000 | 4 |  |
| Flux N2 | 2000 | 4 |  |
| CO2 Low | 1400 | 4 |  |
| CO2 High | 1600 | 5 |  |
| CO2 Ref. | 900 | 4 |  |

| ## Flow | ### Value |
| --- | --- |
| ## Ref CO2 (45) | 45 |
| Flux N2 (45) | 47 |
| Nafion N2 (51) | 47/51 |
| Sample (51) | >60/51 |

| #### Miscellaneous |  |
| --- | --- |
| #### Pressure Gauge | #### 4.5 |
| #### Voltage (from power supply under bench) | #### 15 |
| Current (Amps) | 3 |
| Utility meter (Kilowatts) | 18190 |
| Temperature (shown on heater) | Off |
| Weather | Cloudy and Breezy |

Comments:     Soon after I turned on the computer and opened explorer, the site lost power.  I was able to download the data using the battery back up power and shut the computer down and left it like that when I left.  The data will show a rainfall event for the time I was there, it wasn’t raining but I did unplug the rainbucket and it tipped several times.

Check: pumps running (inside/outside), Licor on and running , UPS operative (under bench, light bulb icon in system tray), PC Anywhere running (small icon of computer in system tray), phone operative (dialtone), breakers not tripped (under bench and at utility meter), shed is dry, tarp covering workbench, heater set to 56, trash emptied, cylinders secured, cabinets and shed locked, rain gauges are clear, tower and guylines up and clear of tree branches, snow cleared at tower base